## Slide 1: Reflectometry Concepts
Hanna Wacklin
Instrument Scientist
European Spallation Source ESS AB Lund, Sweden
MAX IV
&
ESS
Lund
Malmö
Öresund bridge
Copenhagen
KU 2012-08-15

## Slide 2
MAX IV
&
ESS
Lund
Malmö
Öresund bridge
Copenhagen
Reflection from a planar interface
Reflection from thin films
Experiment
Analysis
Reflectometry Concepts

## Slide 3
Reflection from a Planar Interface
The initial neutron with wave vector ki may undergo specular reclection (sr), off-specular reflection (osr), incoherent reflection (icr) or transmission (t), which change its wavevector but not its energy (elastic scattering).
In specular reflection, the reflection angle θf equals the incident angle θi.
For a homogeneous material the potential is the same all over the x,y-plane – consider only z-component of k. (largest application today is disordered surfaces)
Reflection and transmission of neutrons:
ki
kicr
ksr
kosr
kt
i
f
ρ1
ρ2
z
y
x
ρ1≠ ρ2
k = 2π/λ

## Slide 4
Neutron wavefunction
Ψ(r) is the neutron wavefunction
E is the neutron energy – does not change in elastic scattering
V(r) is the nuclear potential of the material which the neutron travels through:
V(r) = is only non-zero when neutron position R = rj (nuclear radius ~10-15m)
-> nuclei scatter like point objects (have no form factor)
Scattering length density in a homogeneous material:
Schrödinger equation for an elastically scattered neutron:
(1)
(2)
(3)
constant potential

## Slide 5
Neutron wavefunction
kx is the wavevector that describes the neutron motion parallel to the interface
– does not change in elastic scattering
Combining (1) and (4):
is the momentum transfer or scattering vector normal to the interface.
The neutron wavefunction is a complex plane wave:
(4)
(5)
where
(6)
(averaging over y)

## Slide 6
Reflection from a Planar Interface
In specular reflection, the reflection angle θf equals the incident angle θi.
For a homogeneous material the potential is the same all over the x,y-plane – consider only z-component of k.
ki
ksr
kt
i
f
ρ1
ρ2
z
y
x
ρ1≠ ρ2
k = 2π/λ
q1
-q1
q2
Δq = 2q1

## Slide 7
Reflection and Transmission coefficients
Where r and t are the reflection and transmission coefficients. The boundary condition is that Ψ(z) and its derivatives are continuous at z=0:
q1and q2 are the scattering vectors above and below the interface. This gives:
The neutrons are partially reflected and transmitted:
(7)
(8)
(9)
and
and
These are the Fresnel reflection and transmission coefficients.
z ≥ 0
z ≤ 0

## Slide 8
Q and Scattering length density
The neutrons are partially reflected and transmitted:
(10)
For elastic scattering:
ki
k1 reflection
k2 transmission
1
1
ρ1
ρ2
z
y
x
2
q1z = k1sin
q2z = k2sin
kx = k1cos1 = k2cos2
Using (6) :
(11)
Δq = Q = 2q1
= 2k1sinθ

## Slide 9
Q and Scattering length density
Using (3), (11) and the de Broglie relation for the neutron energy (h2/2mλ2) gives the neutron refractive index and its relation to the scattering length densities of the two media.
Total internal reflection occurs when θ2 = 0 for transmission and the critical angle is
The refractive index is the ratio between the wavevectors in 2 media:
(12)
(13)
(14)
The critical momentum transfer vector qc depends on ρ1 and ρ2 for any λ.
or
Snell’s Law
(15)

## Slide 10
Reflection from a homogeneous thin film
The neutrons are partially reflected and transmitted:
(16)
The reflection amplitude and intensity are given by:
The intensity maxima and minima occur at cos2β=±1, which allows the film thickness to be estimated from the separation of minima in q:
(17)
ki
k2
0
k1
0
2
0
2
1
1
k1
0
β phase thickness of the film
n is the neutron refractive index
t is the film thickness

## Slide 11
Reflection from a multilayer film
For each layer the reflected amplitude is described by a characteristic Abeles matrix:
For a multilayer film with several sublayers with different scattering length density:
(18)
(19)
The total reflection amplitude is given by the product of all the layers
(20)
0
1
2
3
4
0
1
2
3
r01
r12
r23
r34
t01
t12
t23
t34
The Fresnel reflection coefficient r can be modified to include the effects of interfacial roughness:
(21)
The total reflection intensity is

## Slide 12
The Experiment
Since q = 4πsinθ/λ, reflectivity R(q) can be measured in two ways:
using a monochromatic beam and scanning the angle of incidence θ
using a polychromatic beam covering a range of q-values at any given θ
In practice, method 2 is useful a the range of q is measured simultaneously, which allows structural changes to be monitored as a function of time.

## Slide 13
The Experiment
Since q = 4πsinθ/λ, reflectivity R(q) can be measured in two ways:
using a monochromatic beam and scanning the angle of incidence θ
using a polychromatic beam covering a range of q-values at any given θ
In practice, method 2 is useful a the range of q is measured simultaneously, which allows structural changes to be monitored as a function of time.

## Slide 14
The Experiment
Since q = 4πsinθ/λ, reflectivity R(q) can be measured in two ways:
using a monochromatic beam and scanning the angle of incidence θ
using a polychromatic beam covering a range of q-values at any given θ
In practice, method 2 is useful a the range of q is measured simultaneously, which allows structural changes to be monitored as a function of time.

## Slide 15
The Experiment
Since q = 4πsinθ/λ, reflectivity R(q) can be measured in two ways:
using a monochromatic beam and scanning the angle of incidence θ (x-rays)
using a polychromatic beam covering a range of q-values at any given θ (neutrons)
In practice, method 2 is useful a the range of q is measured simultaneously, which allows structural changes to be monitored as a function of time.
When using polychromatic beam, the neutron wavelength distribution is determined from their time-of-flight, since the neutron velocity depends on their wavelength:
neutron velocity
time-of-flight distance
neutron wavelength

## Slide 16
The Experiment
Reflectivity is presented as the ratio of reflected intensity to the incident beam:
below the critical q, IR/R0= 1 which is used for normalisation to an absolute scale
for a time-of-flight experiment, only 10% of the incident beam is reflected as above the critical angle the intensity falls with ~q-4
Io
IR x 10
I0/ IR
λc ~10.35 Å
log(I0/ IR)
qc ~0.0147 Å-1 (Si –D2O)
The critical q value is a property of the reflecting interface:
for silicon ρ1 = 2.07 x 10-6 Å-2
for D2O ρ2 = 6.35 x 10-6 Å-2
qc (Si-D2O) = 0.0147 Å-1
ρ2 > ρ1  total external reflection

## Slide 17
The Experiment
Neutrons penetrate materials easily, but…they are strongly scattered by liquids (incoherent scattering = contains no structural information), and also absorbed by others.
Many solid materials are almost transparent to neutrons (they don’t absorb neutrons) and can be used to deliver the beam to buried interfaces, e.g. the solid-liquid interface
Io
liquid trough/flow cell
Si (111): 80×50×10 mm
Aluminium holder
Aluminium holder
Silicon crystal
many kinds of sample cells and environments are used for liquid surfaces, e.g.
air-liquid troughs
Langmuir troughs
overflowing cylinder
multi-rack for S/L cells
A typical solid-liquid flow cell:

## Slide 18
Reflection data – what does it reveal?
Neutrons reflection measures the film thickness and scattering length density.
From these is is possible to work out:
volume fraction ϕ or the water fraction (1- ϕ)
the thickness t of a film (or extension length of a molecule from the interface)
the area per molecule A and the surface excess of a film Γ
Using deuterium labeling it is also possible to resolve:
the fraction of a component in a multicomponent film
the location of components in a multicomponent film
the orientation of selectively labeled molecules
volume of water

## Slide 19
Model structure of lipid membrane
= layered model of structure & composition
z
Log R
Neutron wavelength << light wavelength
nanothin films e.g. lipid membranes (~ 5nm)
Solid = Si-SiO2
liquid = D2O
Scattering length density profile (z)
= neutron refractive index profile
SiO2
D2O
Reflection data – what does it reveal?
Optical Matrix modeling

## Slide 20
Scattering length density profile (z)
= neutron refractive index profile
Log R
h-lipid
d-lipid
z
thickness
ρ(z)
ρ
SiO2
D2O
Model structure of lipid membrane
= layered model of structure & composition
z
Solid = Si-SiO2
liquid = D2O
Reflection data – what does it reveal?

## Slide 21
Scattering length density profile (z)
= neutron refractive index profile
Model structure of lipid membrane
= layered model of structure & composition
Log R
SiO2
D2O
h-lipid
d-lipid
z
thickness
ρ(z)
ρ
z
Solid = Si-SiO2
liquid = D2O
Reflection data – what does it reveal?

## Slide 22
Scattering length density profile (z)
= neutron refractive index profile
Model structure of lipid membrane
= layered model of structure & composition
z
Log R
h-lipid
d-lipid
z
thickness
ρ(z)
ρ
z
Solid = Si-SiO2
liquid = D2O
SiO2
D2O
Reflection data – what does it reveal?

## Slide 23
Membrane proteins
5%
50%
11%
50uM
Pore
Membrane lysis by peptides
Kalata peptides: cyclic plant insecticides/Herbal medicine
5uM
5%
25uM
T cell immune receptors and ligands
Collaboration with A. van der Merwe, University of Oxford
Collaboration with D.Craik/C.Wang, UQLD Brisbane

## Slide 24
More applications…
spontaneous adsorption of surfactants, polymers and proteins to interfaces structural changes in response to external parameters: e.g thermo- and pH-responsive polymers and polymer brushes
wetting, dewetting, solvent peneration and drying of films
effect of an interface on organization of ordered mesophases, e.g. liquid crystals, micelles
Some limitations…
surfaces have to be smooth – rms < 5Å desirable
surfaces have to be flat
surfaces need a relatively large area – 5 x 5 cm2 today – at ESS aiming for 5 x 5 mm2
measurement times relatively long – 10s possible today but on large surfaces – at ESS aiming for millisecond-second range
for maximum benefit, materials need to be deuterated – sometimes difficult! setting up a deuteration facility at ESS will be essential.
Information only in surface normal direction (z)

## Slide 25
Off-specular reflection:
What about lateral ordering or correlations on surfaces - 2D structures?
off-specular reflectivity (when q has an x-component) probes structure in the direction of the beam along the surface – e.g. fluctuations. Limited by neutron coherence length (~100μm) – large scale stuctures in e.g. magnetism
The diffraction pattern below TSC consists of a series of sharp peaks as a function of Q. This is characteristic of a stripe-like magnetic domain pattern x in real space.
J.Chakhalian et al. Nature physics, VOL 2, 2006

## Slide 26
Grazing incidence scattering:
What about lateral ordering or correlations on surfaces - 2D structures?
Grazing Incidence Small Angle Scattering (GISAS) probes by x-and y-directions. The difference between GISAS and reflectivity measurements is the beam collimation:
W.Kreutzpaintner et al. Eur. Phys. J. Special Topics 167, 73–79 (2009)
Reflectivity data is measured with a divergent wide beam in the y-direction – this averages lateral structure in the sample plane
GISAS data is measured with a narrow low-divergence beam in the y-direction – this determines the lateral resolution to structure in the sample plane
GISAS

## Slide 27
What about lateral ordering or correlations on surfaces - 2D structures?
Grazing Incidence Small Angle Scattering (GISAS) probes by x-and y-directions – the difference is that the beam needs to be well-collimated in both directions.
With time-of-flight GISANS can probe surface structure at different depths from interface, because at a given angle, the penetration depth increases with decreasing wavelength!
Fig. 1. Selection of twelve 2d intensities measured simultaneously by TOF-GISANS from a nanospolymer nanodot film.
The corresponding mean wavelengths are a) 0.50, b) 0.55, c) 0.61, d) 0.68, e) 0.75, f) 0.82, g) 0.91,
h) 1.00, i) 1.11, j) 1.22, k) 1.35 and l) 1.48 nm. P. Müller-Bushbaum et al. Eur. Phys. J. Special Topics 167, 107–112 (2009)
λ
dp
Grazing incidence scattering:
GISANS can probe buried interfaces that are not accessible by AFM

## Slide 28
Summary
Specular Reflectivity
probes structure in the surface normal direction only – lateral structure is averaged
composition of film, location (and sometimes orientation) of components is obtained from scattering length density profiles
Off-specular Reflectivity
arises when the sample has in-plane structure along the beam direction (x) – resolution limited by neutron coherence length
applications in magnetism, fluctuations, large domain structures
Grazing Incidence Small Angle Scattering (GISAS)
probes by x-and y-directions with pinhole beam
= SANS from the surface only
time-of-flight GISANS probes different depths from interface as function of wavelength
many applications in both soft and hard condensed matter – new technique with analysis still in development
X-rays has higher resolution but neutrons have the isotopic sensitivity
 X-rays for single materials, neutrons for multicomponent systems